## Slide 1
CHEST X-RAY DOMAIN-SHIFT &
SHORTCUT-LEARNING AUDIT
A pneumonia classifier that scores 97% in the hospital it was trained in—and why that number alone should not earn anyone's trust.
Honors Case Study — Baseline CNN, ResNet-50, DenseNet-121 — Kaggle → NIH → OpenI
Notes: Good morning/afternoon. This project audits three pneumonia-detection models -- not just for accuracy, but for why they fail once they leave the hospital they were trained in. I'll walk through the setup, the headline domain-shift result, and what it means for a real deployment decision.

## Slide 2
MOTIVATION
Domain shift is a documented failure, not a hypothetical one
Deep learning models for medical imaging routinely reach high in-distribution accuracy, then fail silently once deployed outside the hospital system they were trained on.
Zech et al., 2018
A pneumonia CNN partly learned to recognize hospital-specific scanner artifacts and text markers — not pathology.
DeGrave et al., 2021
The same shortcut-learning behavior found in COVID-19 chest X-ray classifiers — this is a pattern, not a one-off.
THE QUESTION THIS PROJECT ASKS
Is domain shift an unexplained black box—or a specific, testable mechanism?
This audit trains three architectures on one pediatric hospital's data, tests them—without retraining—on two independent adult sites, and measures whether shortcut reliance tracks a specific, quantifiable cause.
02
Notes: Two published papers set this up. Zech et al. showed a pneumonia CNN partly relying on scanner artifacts instead of pathology. DeGrave et al. found the same shortcut-learning pattern in COVID classifiers. So the real question isn't just whether domain shift is real -- it's whether we can explain it with a specific, testable mechanism instead of shrugging and calling it 'domain shift.'

## Slide 3
RESEARCH QUESTIONS
Five questions, one thread
1
How accurately do CNN architectures detect pneumonia, and what drives misclassification?
2
How does performance degrade under domain shift across two independent external sites?
3
Do a custom CNN, ResNet-50, and DenseNet-121 differ meaningfully in generalization?
4
Is the model attending to lung tissue—or exploiting shortcut features, measured quantitatively?
5
Does shortcut reliance track age group specifically, or a more directly measurable proxy?
03
Notes: Five questions guided this work, but they really boil down to one thread: does the model know what pneumonia looks like, or has it learned a shortcut that happens to correlate with the training hospital? Everything after this slide is in service of answering that.

## Slide 4
DATA & LABEL HARMONIZATION
Three institutions, one binary label, no shared schema
TRAINING
Kaggle
Guangzhou Women & Children's — pediatric
5,856
usable images
73% pneumonia-positive
EXTERNAL VALIDATION
NIH ChestX-ray14
Second institution — adult
61,792
usable images
2.3% pneumonia-positive
EXTERNAL VALIDATION
Indiana Univ. / OpenI
Third institution — adult, public
2,854
usable images
5.5% pneumonia-positive
Patient-level split (not image-level): Kaggle patient IDs parsed from filenames, re-split 4,416 / 560 / 880 with zero patient overlap across train/val/test — verified, not assumed.
Harmonization rule (NIH & OpenI): label = 1 if “pneumonia” is present; label = 0 only if the record is explicitly “normal” / “no finding.” Anything else is an ambiguous negative — excluded, not guessed.
CheXpert (gated, no fixed access timeline) and VinDr-CXR (142GB; its competition packaging has no Pneumonia label) were both investigated and ruled out before settling on OpenI.
04
Notes: Three institutions: Kaggle's pediatric set for training, NIH and Indiana University's OpenI collection for external validation, never trained on. Note the prevalence swing -- 73 percent positive in training versus 2 to 6 percent in the external sets. That mismatch matters a lot later. Also flagging the patient-level split: images from the same patient can't leak across train and test, which the official Kaggle split doesn't guarantee on its own.

## Slide 5
EXPERIMENTAL DESIGN
Three architectures, one evaluation pipeline
Baseline CNN
From scratch · 4 conv blocks · 20 epochs
Lower-bound benchmark
ResNet-50
ImageNet-pretrained · fine-tuned · 10 epochs
Standard transfer baseline
DenseNet-121
ImageNet-pretrained · fine-tuned · 10 epochs
Dense feature reuse
EVALUATION PIPELINE
1
Bootstrapped 95% CIs + McNemar's exact test for architecture differences
2
Three-way domain-shift eval (Kaggle → NIH → OpenI), no retraining between sites
3
Grad-CAM vs. a pretrained lung-segmentation model, full test set (not spot-checked)
4
Border/text-region masking ablation
5
Logistic regression: shortcut reliance ~ age group + resolution + text-marker proxy
05
Notes: Three architectures: a from-scratch baseline CNN, and two ImageNet-pretrained transfer models, ResNet-50 and DenseNet-121. The evaluation pipeline runs statistical testing, three-way domain-shift evaluation, Grad-CAM auditing, a border-masking ablation, and a causal regression -- I'll walk through each of these in turn.

## Slide 6
RESULT · IN-DOMAIN PERFORMANCE
Transfer learning wins — and it's not noise
Model | Accuracy | Precision | Recall | AUC-ROC
Baseline CNN | 0.936 | 0.968 | 0.940 | 0.975
ResNet-50 | 0.972 | 0.970 | 0.990 | 0.992
DenseNet-121 | 0.964 | 0.965 | 0.984 | 0.991
n = 880, held-out Kaggle test set (patient-level, never seen during training)
MCNEMAR'S EXACT TEST
Baseline vs. ResNet-50
p = 1.5×10⁻⁵ — significant
Baseline vs. DenseNet-121
p = 1.8×10⁻⁴ — significant
ResNet-50 vs. DenseNet-121
p = 0.19 — not significant
Transfer learning significantly beats training from scratch. The choice between ResNet-50 and DenseNet-121 mostly doesn't matter.
06
Notes: In-domain, both transfer models clearly beat the baseline, and McNemar's test confirms that's not noise -- p less than 0.001 for both comparisons. Interestingly, ResNet-50 and DenseNet-121 are statistically indistinguishable from each other. So transfer learning matters; which specific pretrained backbone you pick, less so.

## Slide 7
RESULT · THE HEADLINE FINDING
AUC-ROC collapses under domain shift — and not uniformly
Chart: AUC-ROC
| Kaggle (in-domain) | NIH (external) | OpenI (external)
Baseline CNN | 0.975 | 0.62 | 0.512
ResNet-50 | 0.992 | 0.707 | 0.575
DenseNet-121 | 0.991 | 0.679 | 0.591
0.512
Baseline CNN's AUC-ROC on OpenI — 95% CI 0.469–0.556, straddling chance.
Not one “domain shift” effect
Each architecture ranks NIH vs. OpenI differently. A single external site cannot stand in for “externally validated.”
07
Notes: Here's the headline result. All three models collapse when moved to external data, but not the same way -- baseline CNN's AUC-ROC on OpenI is 0.512, with a bootstrapped confidence interval that actually straddles 0.5, so it's not reliably better than chance. ResNet-50 flips the ranking between NIH and OpenI. If domain shift were one monolithic effect, every model would rank the two external sites the same way. They don't -- direct evidence this isn't a single unexplained black-box effect.

## Slide 8
RESULT · QUANTITATIVE SHORTCUT METRIC
Correct answers, wrong reasons — most of the time
12–16%
mean Grad-CAM / lung overlap across all three models
(near-zero would be random; ~100% would be genuine lung-focused attention)
74–81%
of correctly classified cases fall under the 30% overlap cutoff
shortcut reliance is the median case, not a tail risk
BORDER / TEXT-REGION MASKING — ACCURACY DROP WHEN THE OUTER 10% IS BLANKED
Baseline CNN
−14.7 pts (0.913 → 0.767)
ResNet-50
−1.3 pts (0.967 → 0.953)
DenseNet-121
−3.3 pts (0.980 → 0.947)
08
Notes: This is the quantitative core of the shortcut-learning claim. Even on correct predictions, Grad-CAM attention overlaps the segmented lung field only 12 to 16 percent of the time, and 74 to 81 percent of correct predictions fall under our 30 percent shortcut threshold. The border-masking test adds a mechanistic detail: blanking the outer 10 percent of the image costs the baseline CNN almost 15 accuracy points, but barely touches the transfer models -- so the specific shortcut differs by architecture, even though all three are looking somewhere they shouldn't.

## Slide 9
NEGATIVE CASE
Maximum confidence. Maximum wrong.
Truth: NORMAL
Predicted: PNEUMONIA (p = 1.00)
Grad-CAM attention sits almost entirely on the “R” laterality marker in the top-left corner — not on any lung tissue.
This is not a defensible borderline call. It is a textbook shortcut-learning failure, at the model's highest possible confidence — exactly the failure mode this project set out to quantify, not just illustrate.
09
Notes: This single case makes the abstract statistic concrete. Normal chest X-ray, predicted pneumonia with 100 percent confidence -- and the entire Grad-CAM attention sits on the letter R laterality marker in the corner. Maximum confidence, and completely wrong for a reason that has nothing to do with the lungs.

## Slide 10
RESULT · AGE-ARTIFACT CAUSAL ANALYSIS
Is shortcut reliance explained by age—or by resolution?
shortcut_driven ~ age_group + log(image area) + aspect ratio + text-marker proxy
Model | n | age_group | resolution | text marker | overall model
Baseline CNN | 60 | n.s. (p=0.54) | sig. (p=0.04) | n.s. (p=0.52) | —
ResNet-50 | 188 | sig. (p=0.02) | sig. (p=0.01) | sig. (p=0.04) | p=1.4×10⁻¹¹
DenseNet-121 | 194 | n.s. (p=0.36) | n.s. (p=0.11) | n.s. (p=0.28) | p=6.1×10⁻¹⁸
TAKEAWAY
“Age group” is not a consistent, well-explained predictor of shortcut reliance across architectures. Where a factor is significant, it's resolution and/or detected artifacts — measurable, mechanistic proxies — not an unexplained institutional label. That's the difference between a black-box excuse and a testable causal claim.
10
Notes: We asked: is this shortcut reliance actually about pediatric-versus-adult age, or something more mechanical, like image resolution? The answer is mixed across architectures, which is itself informative. For the baseline CNN, age isn't significant once you control for resolution -- resolution alone explains it. For ResNet-50, all three factors are significant, including the text-marker proxy, which validates that our coarse heuristic is picking up something real. The one consistent finding across all three: age group by itself is not a reliable explanation for shortcut reliance.

## Slide 11
RESPONSIBLE AI & CLINICAL DEPLOYMENT
What these numbers mean for a real deployment decision
Bias & data quality
Single-institution, pediatric training data — required active correction: patient-level re-split and re-download of ~12% corrupted training images found during verification.
Shortcut reliance is the median case
74–81% of correct predictions still show below-threshold lung overlap. This isn't a tail-risk QA catch — it's how these models mostly reach their answers.
Generalization isn't one number
Two independent external sites degrade differently per architecture. Single-site “external validation” is not sufficient grounds for deployment claims.
Deployment recommendation
None of the three models should run autonomously outside a population resembling the training data. At most: a low-stakes second-reader signal, mandatory radiologist review, site-specific threshold recalibration, and ongoing shortcut-overlap monitoring.
11
Notes: Pulling this together for a deployment conversation: the training data itself needed active correction -- we found and fixed corrupted files and patient-level leakage. Shortcut reliance isn't rare, it's the majority case. Generalization isn't a single number since the two external sites disagree. And the bottom line: none of these three models should run autonomously outside a population like the training data.

## Slide 12
LIMITATIONS
Read these numbers as real, not as ceilings
Training budget (20 / 10 / 10 epochs) demonstrates real, measured phenomena — not necessarily converged optimal performance.
Text-marker detection is a coarse OpenCV heuristic, not OCR — no OCR engine was available in the build environment.
Causal-analysis samples (60–194 per model) are modest for a four-predictor logistic regression; individual coefficients carry wide intervals.
Two external sites, not the originally planned three — CheXpert's access gate and VinDr-CXR's label/size mismatch made a third impractical in scope.
The false-negative failure taxonomy is a heuristic first-pass triage, not a validated clinical categorization.
12
Notes: A few honesty notes before the conclusion: these are real but not necessarily fully converged models. Text-marker detection is a coarse heuristic, not real OCR, since no OCR engine was available. The causal-analysis samples are modest, so individual coefficients should be read cautiously. And we only got two of the three originally planned external sites, since CheXpert and VinDr-CXR both turned out to be impractical for documented reasons, not just skipped.

## Slide 13
CONCLUSION
Transfer learning wins in-domain. Neither transfer model earns trust out-of-domain.
Both transfer architectures significantly beat the from-scratch baseline in-domain (McNemar's p < 0.001) — but that edge does not carry over to external generalization.
All three models show substantial, non-uniform degradation on independent adult-population sites — partly threshold miscalibration, partly measured reliance on non-lung image regions.
Where shortcut reliance is statistically explained at all, it tracks resolution and detected artifacts specifically — not an unexplained “domain shift” or “age” label.
A testable causal mechanism, not a black-box excuse — and a concrete argument against single-site validation as grounds for clinical deployment.
13
Notes: So: transfer learning wins in-domain, decisively. But that in-domain edge doesn't buy trustworthy generalization -- all three models fail non-uniformly outside their training distribution, and where we can explain that failure statistically, it's resolution and detected artifacts, not an unexplained domain-shift or age label. That's the difference between a black-box excuse and a testable, actionable finding -- and a concrete argument against single-site validation as grounds for clinical deployment. Thank you.